# ReportScreen.kt — Android ECG Report Screen & PDF Renderer

File: ReportScreen.kt — 1181 lines

Package: com.deckmount.ecgapp.presentation.ui.screen.report

Platform: Android, Jetpack Compose + Material 3

Original author: Amarjeet Kumar

Last revised: 2026-08-18

# 1. Purpose

This file is the Android counterpart to the desktop report generators in

src/ecg/. It renders a finished 12-lead ECG report — grid, header, waveforms,

calibration pulses, conclusions and footer — and does so through **one drawing implementation that

serves both the on-screen preview and the exported PDF**.

That single-implementation property is the central design decision of the file. Every coordinate is

expressed in millimetres, and a single scale factor (pxPerMm) converts millimetres to whatever

the current output device wants:

| Output | pxPerMm | Origin |
| --- | --- | --- |
| Screen preview | (availableWidthPx / pageWidthMm) × 2.5 | Composable measures itself, then oversamples 2.5× |
| PDF | 595 / 210 ≈ 2.833 (portrait) or 842 / 297 ≈ 2.835 (landscape) | A4 in PDF points, 1 pt = 1/72 in |

Because both paths run ECGReportRenderer.draw(canvas), what the clinician sees on the phone is what

prints on paper. There is no second layout engine to drift out of sync.

# 2. File map

| Lines | Element | Kind |
| --- | --- | --- |
| 1–56 | Package and imports | — |
| 62–95 | Page, margin, signal and grid constants | private const val |
| 101–330 | ECGReportScreen | @Composable — public entry point |
| 337–441 | ECGReportCanvas | @Composable — private preview surface |
| 449–1130 | ECGReportRenderer | internal class — all drawing |
| 1141–1182 | generateECGPdf | top-level function |

Within ECGReportRenderer:

| Lines | Member |
| --- | --- |
| 454–478 | Derived geometry, isForPdf, p() / pt() unit helpers |
| 481–549 | Paint factory and pre-built paints |
| 556–562 | drawTxt — TCPDF-compatible text placement |
| 580–595 | draw — the top-level draw order |
| 602–618 | drawGrid |
| 624–721 | drawHeader |
| 729–756 | draw1x12 — portrait |
| 766–820 | draw2x6 — landscape |
| 829–885 | draw3x4 — landscape |
| 891–994 | drawFooter / portrait / landscape |
| 1016–1049 | drawCalibration, drawCalibrationPad |
| 1096–1129 | drawWaveform |

# 3. External contract

## 3.1 ECGReportRenderData

Defined outside this file (in the sibling package

com.deckmount.ecgapp.presentation.ui.screen.livemonitorecg). Every field this file reads:

| Field | Type | Used for |
| --- | --- | --- |
| layout | String | Layout selector: "1x12", "3x4", anything else → "2x6" |
| patientName, patientAge, patientGender | String | Header column 1; blank renders as - |
| reportDate, reportTime | String | Header column 1, pre-formatted — this file does no date formatting |
| acFilter | String/numeric | Printed into the spec line as AC:<value>Hz |
| hr, rr, pr, qrs, qt | numeric | Header column 2 |
| qtc, qtcf | numeric | Header column 3 |
| rv5, sv1 | Float | Header column 3; printed to 3 dp, summed for the Sokolow-Lyon index |
| orgName, orgAddress, orgPhoneNo | String | Header right block; phone is prefixed +91 when non-blank |
| machineSerial | String | Footer — last 4 characters only |
| conclusions | List<String> | Footer conclusion box |
| leadData3500 | Map<String, List<Float>> | 1×12 portrait strips |
| leadData1750 | Map<String, List<Float>> | 2×6 landscape strips |
| leadData1250 | Map<String, List<Float>> | 3×4 landscape strips |
| leadData5000 | Map<String, List<Float>> | Rhythm strip (lead II) in both landscape layouts |

pqrstAxis exists on the model but is referenced only inside a commented-out line

(ReportScreen.kt:687) — the P/QRS/T axis is not printed today.

A missing map key degrades gracefully: samples[lead] ?: emptyList(), and drawWaveform returns

immediately when samples.size < 2. A missing lead therefore prints as an empty lane with its label and

calibration pulse intact — no crash, no exception.

## 3.2 ALL_LEADS

An external List<String> imported from the live-monitor package, iterated by draw1x12. The 2×6 and

3×4 layouts do not use it; they carry their own hardcoded lead orderings (§7.2, §7.3).

## 3.3 Sample-count contract

| Map | Samples | Duration @ 500 Hz | Width at 25 mm/s | Lane width available | Headroom |
| --- | --- | --- | --- | --- | --- |
| leadData3500 | 3500 | 7.0 s | 175.0 mm | 185 mm | 10 mm |
| leadData1750 | 1750 | 3.5 s | 87.5 mm | 123 mm | 35.5 mm |
| leadData1250 | 1250 | 2.5 s | 62.5 mm | 80 mm | 17.5 mm |
| leadData5000 | 5000 | 10.0 s | 250.0 mm | 254 mm | 4 mm |

Every stream fits its lane. Anything longer is silently truncated — drawWaveform breaks out of its

loop the moment xMm > maxXMm (ReportScreen.kt:1120) with no warning and no

log. If a caller ever supplies more samples than the table above, the tail simply disappears from the

report. Treat these counts as a hard contract.

# 4. Constants

private const val A4_P_W = 210f     // portrait page width  (mm)
private const val A4_P_H = 297f     // portrait page height (mm)
private const val A4_L_W = 297f     // landscape page width  (mm)
private const val A4_L_H = 210f     // landscape page height (mm)

private const val MARGIN_PDF    = 0f    // ⚠ declared, never referenced
private const val MARGIN_TOP    = 5f
private const val MARGIN_BOTTOM = 5f
private const val MARGIN_LEFT   = 5f
private const val MARGIN_RIGHT  = 5f

private const val ADC_PER_MV = 1.8f     // ADC units per millivolt
private const val ECG_FS     = 500f     // sampling rate (Hz)

private const val FIXED_WAVE_SPEED = 25f    // mm/s  — not user-adjustable
private const val FIXED_WAVE_GAIN  = 10f    // mm/mV — not user-adjustable

private const val GRID_W_PORTRAIT  = 210f
private const val GRID_H_PORTRAIT  = 297f
private const val GRID_W_LANDSCAPE = 297f
private const val GRID_H_LANDSCAPE = 210f

Stale comments. The margin constants are 5f but their trailing comments still say *"10mm"*, and

the layout functions carry header comments citing topOffset=38 / startY=35 from when the margins

were 10 mm. The code is authoritative; the derived values in §7 are computed from 5f. The older

copy of this file at test_cloud_connection/ReportScreen.kt

still uses 10 mm margins and an extra ADC_PER_MM = 2.75f constant — see §12.

Paper-speed identities (worth internalising, they explain every magic number below):

mmPerSample = FIXED_WAVE_SPEED / ECG_FS = 25 / 500 = 0.05 mm/sample

1 mm ≡ 20 samples ≡ 40 ms

One large grid box (5 mm) ≡ 200 ms horizontally, ≡ 0.5 mV vertically at 10 mm/mV

1 mV ≡ 1.8 ADC units ≡ 10 mm on paper

# 5. Coordinate system

Everything is authored in millimetres and converted at draw time:

private fun p(mm: Float)      = mm * pxPerMm            // mm → device units
private fun pt(ptSize: Float) = ptSize * 0.352778f * pxPerMm   // font pt → device units

0.352778 is mm-per-point (1/72 inch). pt() exists so font sizes can be specified in the same

typographic points the legacy TCPDF templates used, and land at the same physical size on both outputs.

## 5.1 drawTxt — TCPDF text semantics

Android's Canvas.drawText positions text by its baseline. TCPDF's Text($x, $y, $s) positions by

the top-left of the bounding box. Rather than translate every legacy coordinate, the file adapts the

API:

val fm = paint.fontMetrics
canvas.drawText(text, p(xMm), p(yMm) - fm.ascent, paint)   // fm.ascent is negative

Consequence: every yMm passed to drawTxt is the top of the text, not the baseline. This is

the single most common source of confusion when editing layout coordinates in this file.

## 5.2 isForPdf

private val isForPdf = pxPerMm < 3.5f

An inferred flag, not a parameter. PDF rendering runs at ≈ 2.83 units/mm; screen rendering at

availableWidthPx / pageWidthMm × 2.5, which on any normal device is 9–13. It selects thinner strokes

for print (§6).

⚠️ Fragile heuristic. The flag flips to true whenever the composable is narrower than

≈ 294 px (portrait) or ≈ 415 px (landscape) — reachable in split-screen, a foldable cover display, or a

small preview pane. The screen preview would then draw with print-weight hairlines. Passing an explicit

isForPdf: Boolean into the constructor would remove the guesswork entirely.

# 6. Paints

Built once per ECGReportRenderer instance, all anti-aliased.

| Paint | Colour | Stroke (mm) | Purpose |
| --- | --- | --- | --- |
| gridMinor | RGB(245, 220, 220) | 0.1 | 1 mm grid |
| gridMajor | RGB(230, 150, 150) | 0.25 | 5 mm grid |
| waveP | Black | PDF 0.20 / screen 0.25 (min 1.5 px) | ECG trace — round join and cap |
| calibP | Black | 0.4 (min 1.5 px on screen) | 1 mV calibration pulse — mitre join, square cap |
| boxP | Black | 0.3 (min 1.0 px on screen) | Conclusion box outline |
| dashDivP | RGB(80, 80, 80) | 0.4 (min 1.2 px on screen) | Dashed 2 mm/2 mm column divider |

The coerceAtLeast floors exist because a 0.25 mm line at low screen density can round to zero device

pixels and vanish; the floor guarantees the waveform is always visible on screen. PDF has no floor

because vector output has no minimum stroke.

Text paints are produced by mkText(ptSize, bold, italic) and cached as tp6…tp20_5B. The naming is

positional: tp10_5B = 10.5 pt bold.

Dead paints: tp6 and tp20_5B are declared but never drawn with; tp8B is used only inside a

commented-out logo-fallback block.

# 7. Layouts

draw() establishes the order — background, grid, header, layout body, footer:

fun draw(canvas: Canvas) {
    canvas.drawColor(GColor.WHITE)
    drawGrid(canvas, 0f, 0f, gridW, gridH)   // grid spans the FULL page, margins included
    drawHeader(canvas)
    when (data.layout) {
        "1x12" -> draw1x12(canvas)
        "3x4"  -> draw3x4(canvas)
        else   -> draw2x6(canvas)            // default
    }
    drawFooter(canvas)
}

drawGrid walks integer millimetres across the page (for (i in 0..w.toInt())), drawing a major line

every 5th and a minor line otherwise. There is no grid clipping to a content area — the pink grid runs

edge to edge, exactly as ECG paper does. The inline comment claiming the grid starts at a 5 mm margin is

stale; drawGrid is called at (0, 0).

## 7.1 Header — shared by all three layouts

yBase = MARGIN_TOP - 3f = 2f, line height lh = 5f. Left column starts at x = 5 (portrait) or

x = 10 (landscape).

| Column | Portrait X | Landscape X | Rows (top-of-text, mm) |
| --- | --- | --- | --- |
| 1 — Patient | 5 | 10 | Name 2 · Age 7 · Gender 12 · ECG Type: Standard 17 · Date & Time 22 · spec line 27 |
| 2 — Intervals | 60 | 95 | HR 2 · RR 7 · PR 12 · QRS 17 · QT 22 |
| 3 — Derived | 90 | 125 | QTc 2 · QTcF 7 · RV5/SV1 12 · RV5+SV1 index 17 |
| 4 — Organisation | 140 | 220 | orgName 2 · orgAddress 7 · +91 <phone> 12 |

Two details worth knowing:

Spec line. "25.0 mm/s   0.5-25 Hz   AC:<acFilter>Hz   10.0 mm/mV". Speed and gain are

interpolated from the constants, so they can never disagree with what was actually drawn. Only the AC

filter comes from the data.

Sokolow-Lyon flag. RV5 + SV1 ≥ 3.5 mV appends a * to the index line

(ReportScreen.kt:677) — a possible-LVH marker. There is no legend explaining

the asterisk anywhere on the page.

The header occupies through y = 27 mm. The landscape body starts at y = 30 mm and the portrait body at

y = 33 mm, leaving 3 mm of clearance in landscape. A longer organisation name does not wrap — it runs

straight into whatever sits to its right.

## 7.2 draw1x12 — portrait, 12 stacked leads

| Quantity | Expression | Value |
| --- | --- | --- |
| topOffset | MARGIN_TOP + 28 | 33 mm |
| usableH | 297 − 33 − 5 − 12 | 247 mm |
| cellH | usableH / 12 | ≈ 20.583 mm |
| Lane centre, lead *i* | topOffset + i × cellH + cellH/2 | 43.3, 63.9, 84.5 … |
| Label | x = 10, y = midY − 10 | above the trace |
| Calibration | x = MARGIN_LEFT = 5 | no pad |
| Waveform | x0 = 18, width = 210 − 5 − 5 − 15 | 185 mm |

Lead order comes from the external ALL_LEADS. Source data: leadData3500 (7 s per lead).

## 7.3 draw2x6 — landscape, six rows of two leads plus a rhythm strip

| Quantity | Expression | Value |
| --- | --- | --- |
| startY | MARGIN_TOP + 25 | 30 mm |
| rowH | — | 22 mm |
| leftMargin | MARGIN_LEFT + 8 | 13 mm |
| Left label / waveform | 22 / 27, width 123 | ends at 150 mm |
| Dashed divider | 13 + 14 + 123 + 5 | 155 mm |
| Right label / waveform | 160 / 165, width 123 | ends at 288 mm |
| Rhythm strip centre | 30 + 6×22 + 2 + 7.5 | 171.5 mm |
| Rhythm waveform | x0 = 27, width 297 − 13 − 5 − 25 | 254 mm |

Pairing is fixed and not the conventional Cabrera or standard column order:

I / V1    II / V2    III / V3    aVR / V4    aVF / V5    aVL / V6

Note the third column: aVR, aVF, aVL — aVF before aVL. The standard augmented-lead order is

aVR, aVL, aVF. If this is intentional it deserves a comment in the source; if not, it is a report defect

and clinicians will notice.

Row leads use leadData1750 (3.5 s); the rhythm strip uses leadData5000["II"] (10 s).

## 7.4 draw3x4 — landscape, four rows of three leads plus a rhythm strip

| Quantity | Expression | Value |
| --- | --- | --- |
| startY | MARGIN_TOP + 25 | 30 mm |
| rowH | — | 30 mm |
| leftPad | MARGIN_LEFT + 8 + 10 | 23 mm |
| leadW | — | 80 mm |
| Column X | 23 + c × 90 | 23, 113, 203 (ends 283 mm) |
| Row centres | 30 + r × 30 + 15 | 45, 75, 105, 135 mm |
| Rhythm strip centre | 30 + 4×30 + 3 + 7.5 | 160.5 mm |
| Rhythm waveform | x0 = 27, width 254 | — |

Groups here are in standard order:

I   II   III
aVR aVL  aVF
V1  V2   V3
V4  V5   V6

Dashed dividers are drawn after columns 0 and 1 only. Row leads use leadData1250 (2.5 s).

## 7.5 Footer

Selected by orientation. Both variants draw a doctor block on the left, a bordered CONCLUSION box on

the right, and a centred device/compliance line.

|  | Portrait | Landscape |
| --- | --- | --- |
| Doctor block top | y = 277 | y = 192 |
| Conclusion box | x 95, y 272, 105 × 18 mm | x 140, y 185, 145 × 20 mm |
| Box title | CONCLUSION, 7 pt bold, centred | CONCLUSION, 9 pt bold, centred |
| Conclusion grid | 3 columns, 3.5 mm row pitch | 3 columns, 5 mm row pitch, 5 mm column gap |
| Centred footer line | y = 292.5, 7 pt | y = 205.5, 8 pt |

Footer text: Deckmount Electronics Pvt Ltd | Rhythm Ultra Max | IEC 60601 | <last 4 of serial> | Made in India,

horizontally centred by measuring the string and dividing the slack.

Two behaviours in the conclusion box are easy to miss:

Zero-HR override. When data.hr <= 0 the list is replaced wholesale with

listOf("No ECG data available") — the app must never print an interpretation for a dead trace. This

mirrors the desktop RPT-13 rule. Otherwise blanks and "---" placeholders are filtered out.

Silent overflow. Items whose row would exceed the box height are skipped

(:935, :985). Since ty grows monotonically, the

first skip effectively caps the list — but the loop keeps iterating, and the portrait variant numbers

items by loop index i while the landscape variant uses a separate counter sr that only advances on

drawn items. The two orientations therefore number a truncated list differently. With three columns

and the current box heights, roughly 9 (portrait) / 9 (landscape) items fit; the desktop pipeline caps

conclusions at 5, so overflow should not occur in practice — but nothing in this file enforces that.

# 8. Calibration pulse

┌───────────┐   ← y − gain (10 mm above baseline)
        │           │
 ───────┘           └───────   ← baseline y
 x    x+2         x+7      x+9

Five line segments, 9 mm wide, 10 mm tall — a standard 1 mV / 0.2 s reference. At 25 mm/s the 5 mm top

segment is exactly 200 ms, so the pulse also reads as one large grid box wide.

drawCalibrationPad(x, …) simply calls drawCalibration(x + 4f, …); the landscape layouts pass

leftMargin − 4f, so the 4 mm pad cancels out and the pulse begins at leftMargin exactly. The

indirection is a leftover from the earlier port and could be removed.

# 9. Waveform rendering

private fun drawWaveform(canvas, samples, x0Mm, y0Mm, widthMm, gainMmPerMv) {
    if (samples.size < 2) return

    val sorted   = samples.sorted()
    val baseline = sorted[sorted.size / 2]                       // median = robust DC estimate
    val scale    = (1f / ADC_PER_MV) * gainMmPerMv * pxPerMm     // ADC → device units

    var prevX = p(x0Mm)
    var prevY = p(y0Mm) - (samples[0] - baseline) * scale

    for (i in 1 until samples.size) {
        val xMm = x0Mm + i * mmPerSample
        if (xMm > x0Mm + widthMm) break                          // clip, silently
        val x = p(xMm)
        val y = p(y0Mm) - (samples[i] - baseline) * scale
        canvas.drawLine(prevX, prevY, x, y, waveP)
        prevX = x; prevY = y
    }
}

Vertical scale. (1/1.8) × 10 × pxPerMm ≈ 5.556 × pxPerMm. One ADC unit is 1/1.8 mV, and at

10 mm/mV that is 5.556 mm — so a 1.8-ADC deflection draws exactly 10 mm, one large box per millivolt.

Y is subtracted because canvas Y grows downward while ECG amplitude grows upward.

Baseline removal. The median of the strip is subtracted, not the mean. On an ECG the median sits

on the isoelectric line because most samples are baseline; the mean would be dragged upward by R waves.

This centres each lane independently and removes per-lead DC offset.

Two consequences worth stating plainly:

Baseline is computed per strip, so the same lead can sit at slightly different vertical offsets in

the 2×6 body (1750 samples) and in the 10 s rhythm strip (5000 samples).

The method is robust to noise but not to *sustained* pathology: a strip where elevation persists across

more than half the samples has that elevation partially normalised away. For the 2.5–10 s windows used

here this is not a practical concern, but it is the reason the desktop pipeline does its own DC removal

upstream rather than relying on the renderer.

No decimation. Every sample becomes one drawLine call. A portrait report is

12 × 3500 ≈ 42,000 line calls; a landscape 2×6 report is 12 × 1750 + 5000 = 26,000. At the 2.5×

oversampled preview size that is many samples per device pixel, so the cost buys no visible detail. The

desktop app solved the equivalent problem by decimating 2× (see the hyperkalemia display notes in the

README); the same optimisation is available here — reduce the sample stride until roughly

2 points land per device pixel column.

No NaN guard. Float.NaN in the sample list propagates into drawLine, which silently drops the

segment. Unlike the desktop connect='finite' path there is no visible break, but there is a gap. The

caller is responsible for sanitising.

# 10. Compose layer

## 10.1 ECGReportScreen(reportData, onNavigateBack)

A Material 3 Scaffold. The top bar carries a back button, a Share action and a Download action;

the body is ECGReportCanvas.

Share flow (:228–263):

shareMutex.tryLock() — returns immediately if a share is already running.

isSharing = true, spinner replaces the icon, and both actions disable

(enabled = !isSharing && !isExporting).

generateECGPdf runs on Dispatchers.IO.

FileProvider.getUriForFile(context, "${packageName}.fileprovider", file) produces a content URI.

Every activity resolving the ACTION_SEND intent is granted FLAG_GRANT_READ_URI_PERMISSION

explicitly, then Intent.createChooser is launched.

finally clears isSharing and unlocks — so a failure cannot wedge the button.

Download flow (:266–308):

Generate the PDF on Dispatchers.IO into the cache directory.

Hold it in tempPdfFile, then launch the Storage Access Framework

(ActivityResultContracts.CreateDocument("application/pdf")) with a suggested name

RhythmPro_ECG_<yyyyMMdd_HHmmss>.pdf.

On a returned URI, stream the temp file into it and delete the temp file, then toast success.

On cancellation, delete the temp file and clear state.

Using SAF rather than a hardcoded path is the right choice — it needs no storage permission on any API

level and lets the user pick the destination.

⚠️ `isExporting` is never set to `true`. Both assignments are commented out

(:270, :285). The download spinner therefore never

appears, and the guard if (isExporting || isSharing) return@IconButton cannot block a second tap

during generation. Rapid double-tapping Download generates two PDFs and opens two file pickers. Share

is protected by its mutex; Download is not. Uncommenting both lines restores the intended behaviour.

## 10.2 ECGReportCanvas(reportData, modifier)

BoxWithConstraints
  └─ pxPerMm      = availableWidthPx / pageWidthMm
     displayW/H   = the exact on-screen page rectangle
     bmpW/H       = display × 2.5          ← quality oversample
     renderPxPerMm= pxPerMm × 2.5

The report is rasterised once into an ARGB_8888 bitmap inside produceState keyed on

(reportData, bmpW, bmpH), on Dispatchers.Default. Pan and pinch-zoom then transform that finished

bitmap through a graphicsLayer (scale clamped to 0.5×–8×, origin centred). Redrawing vector content on

every gesture frame would be far too slow; this trades memory for a smooth gesture.

While the bitmap is null a centred spinner and *"Rendering ECG report…"* are shown. A render exception is

caught, logged under tag ECGReportScreen, and leaves the state null — the spinner then never resolves,

which is the one failure mode with no user-visible explanation.

⚠️ Bitmap memory. ARGB_8888 at 2.5× in both axes is 6.25× the pixel count of the displayed page.

On a 1080 px-wide phone rendering portrait A4: 2700 × 3819 × 4 bytes ≈ 41 MB for a single bitmap,

and the previous one is only freed once produceState reassigns. On a low-RAM device this is a

plausible OutOfMemoryError. Two mitigations, either sufficient: use Bitmap.Config.RGB_565 (the

report has no transparency — halves the cost), or lower qualityMultiplier to 1.5–2.0 and let the

zoom gesture re-render at higher fidelity only when the user actually zooms in.

⚠️ Unbounded pan. offsetX/offsetY accumulate raw gesture deltas with no clamping

(:396–397), so the report can be flung entirely off-screen with no way back

except leaving and re-entering the screen. Clamping to the scaled bitmap bounds — or a double-tap

reset — would fix it.

# 11. generateECGPdf(context, data): File?

val pageWPt = if (isPortrait) 595 else 842      // A4 in PDF points
val pageHPt = if (isPortrait) 842 else 595
val pxPerMm = pageWPt / pageWMm                  // ≈ 2.833 pt/mm

val document = PdfDocument()
val page     = document.startPage(PdfDocument.PageInfo.Builder(pageWPt, pageHPt, 1).create())
ECGReportRenderer(data, pxPerMm, context).draw(page.canvas)
document.finishPage(page)

Output lands in context.cacheDir/shared_pdfs/RhythmPro_ECG_<yyyyMMdd_HHmmss>.pdf. The cache directory is

deliberate: it is the standard FileProvider root for sharing, and Android may reclaim it under storage

pressure. Returns null on any exception, logged at ERROR under ECGReportScreen — every caller

null-checks and toasts *"PDF generation failed"*.

Because pxPerMm is derived from the page size in points, the renderer emits true A4 geometry: a 5 mm

grid box measures 5 mm on paper when printed at 100 % scale.

Housekeeping gap. The Share path writes a fresh PDF into shared_pdfs/ on every tap and never

deletes it (only the Download path cleans up, and only its own temp file). Repeated sharing grows the

cache monotonically until Android reclaims it. Pruning files older than a few hours on entry to the

screen would be a two-line fix.

`SimpleDateFormat` + `Locale.getDefault()`. Under a non-Gregorian default calendar (e.g. a

Thai-locale device) yyyy yields Buddhist-era years, producing filenames like RhythmPro_ECG_2569….

The desktop doctor-review backend parses device id and timestamp out of the filename, so a

non-Gregorian year would break that matching. Locale.US is the correct choice for machine-readable

timestamps.

# 12. Divergence from the sibling copy

test_cloud_connection/ReportScreen.kt is an older copy of

this file. Known differences:

|  | Root ReportScreen.kt | test_cloud_connection/ copy |
| --- | --- | --- |
| Margins | 5 mm | 10 mm |
| ADC_PER_MM | absent | 2.75f |
| Grid size constants | present (GRID_W_* / GRID_H_*) | absent |
| ALL_LEADS / ECGReportRenderData | resolved implicitly | imported explicitly from livemonitorecg |
| HRV imports | present (unused) | absent |

Treat the root file as current. The duplicate is a snapshot and should either be deleted or clearly

marked, because a future edit applied to the wrong copy will silently do nothing.

# 13. Dead code and unused imports

Confirmed by occurrence count across the file:

| Symbol | Status |
| --- | --- |
| adcScaleFactor (:466) | Declared and computed, never read. drawWaveform uses ADC_PER_MV directly. Leaving it in place invites someone to "fix" the scale by editing a constant that does nothing. |
| MARGIN_PDF | Declared, never referenced |
| tp6, tp20_5B | Paints declared, never used |
| tp8B | Used only in a commented-out logo fallback |
| HRVAnalysisResult, HRVReportRenderer, EcgPdfGenerator | Imported, never used |
| kotlinx.coroutines.flow.first | Imported, never used |
| BitmapFactory, Rect, RectF, R | Imported for the commented-out logo block only |
| data.pqrstAxis | Referenced only inside a comment |
| Commented draw() variant (:568) | Superseded by the grid-constant version |
| Commented drawCalibration variant (:1030) | "Left bottom line removed" experiment |
| Commented drawWaveform variant (:1063) | Pre-median-baseline version |

None of it is harmful today. All of it costs a future reader time, and the adcScaleFactor case is an

active trap.

# 14. Threading

| Work | Dispatcher | Why |
| --- | --- | --- |
| Bitmap rasterisation | Dispatchers.Default | CPU-bound drawing; produceState cancels it automatically when the composable leaves |
| PDF generation | Dispatchers.IO | Renders and writes a file |
| Copying the PDF into the SAF URI | Dispatchers.IO | Stream copy |
| Toasts, UI state | Main | Explicit withContext(Dispatchers.Main) inside IO blocks |

Nothing renders on the main thread. shareMutex prevents concurrent share operations; see §10.1 for the

missing equivalent on Download.

# 15. Making changes safely

Changing paper speed or gain. FIXED_WAVE_SPEED and FIXED_WAVE_GAIN are the only inputs to

mmPerSample and to the waveform scale, and both are interpolated into the printed spec line — so the

header can never lie about what was drawn. But the sample-count table in §3.3 assumes 25 mm/s: at

50 mm/s, leadData3500 would need 370 mm and be clipped to just over half. **Changing the speed requires

re-deriving §3.3 and re-checking every lane width.**

Adding a layout. Add the branch in draw(), write a drawNxM following the existing shape

(centre-line per lane → drawCalibrationPad → drawTxt label at midY − 10 → drawWaveform), pick the

matching leadDataN map, and confirm samples × 0.05 mm ≤ laneWidth.

Moving header or footer text. Remember drawTxt takes the top of the text, not the baseline

(§5.1). Verify in both orientations — the footers are separate functions with independent coordinates,

and it is easy to fix one and leave the other.

Before merging any change to this file, export a PDF and print it at 100 % scale, then measure a

large grid box with a ruler. It must be 5 mm. That one measurement catches almost every unit-conversion

regression this file can suffer.

# 16. Verification checklist

Mirrors the format of the desktop EXE test checklist. Prefix KRP.

Rendering

☐  KRP-01 All three layouts render: "1x12", "2x6", "3x4".

☐  KRP-02 An unrecognised layout string falls back to "2x6" rather than crashing.

☐  KRP-03 Grid is 1 mm minor / 5 mm major and spans the full page edge to edge.

☐  KRP-04 Every lane shows its calibration pulse: 9 mm wide, 10 mm tall.

☐  KRP-05 Every lane shows its lead label above the trace.

☐  KRP-06 A missing lead key renders an empty lane — label and calibration intact, no crash.

☐  KRP-07 A single-sample lead list renders nothing and does not throw.

☐  KRP-08 The 2×6 rhythm strip uses lead II from leadData5000 and spans the full 10 s.

Geometry — the checks that need a ruler

☐  KRP-09 Exported PDF printed at 100 % scale: one large grid box measures 5.0 mm.

☐  KRP-10 The calibration pulse measures 10 mm tall and 5 mm across its top segment.

☐  KRP-11 A 1 mV deflection in the trace measures 10 mm — one large box.

☐  KRP-12 One second of signal spans 25 mm — five large boxes.

☐  KRP-13 No strip is clipped: the last sample of each stream is visible within its lane.

☐  KRP-14 Nothing overflows the page in either orientation, including a long organisation name.

Header and footer

☐  KRP-15 Blank patient fields render as -, not as an empty gap or null.

☐  KRP-16 The spec line reads 25.0 mm/s   0.5-25 Hz   AC:<n>Hz   10.0 mm/mV.

☐  KRP-17 RV5, SV1 and the index print to 3 decimal places.

☐  KRP-18 RV5 + SV1 ≥ 3.5 appends *; below the threshold it does not.

☐  KRP-19 The footer shows the last 4 characters of machineSerial and is horizontally centred.

☐  KRP-20 hr <= 0 prints exactly "No ECG data available" in the conclusion box and nothing else.

☐  KRP-21 Blank conclusions and "---" are filtered out.

☐  KRP-22 Conclusions stay inside the box border in both orientations; numbering is sequential from 1.

Interaction

☐  KRP-23 The spinner and *"Rendering ECG report…"* appear, then resolve to the report.

☐  KRP-24 Pinch-zoom clamps at 0.5× and 8×.

☐  KRP-25 Pan moves the report smoothly. *(Known issue: pan is unbounded — record the behaviour.)*

☐  KRP-26 Rotating the device re-renders at the new width without crashing.

☐  KRP-27 Back navigation leaves the screen cleanly and cancels any in-flight render.

Share and download

☐  KRP-28 Share opens the system chooser; the recipient app can open the PDF.

☐  KRP-29 The share spinner shows during generation and both actions disable.

☐  KRP-30 Rapid double-tap on Share starts exactly one operation (mutex).

☐  KRP-31 Rapid double-tap on Download starts exactly one operation.

*(Known issue: currently fails — isExporting is never set. See §10.1.)*

☐  KRP-32 Download opens the SAF picker pre-filled with RhythmPro_ECG_<timestamp>.pdf.

☐  KRP-33 Saving writes a valid PDF to the chosen location and toasts success.

☐  KRP-34 Cancelling the picker deletes the temp file and leaves no orphan in cacheDir/shared_pdfs.

☐  KRP-35 The exported PDF opens in Drive, Adobe Reader and Chrome.

☐  KRP-36 Filenames use Gregorian years on a device with a non-Gregorian default locale.

*(Known issue — see §11.)*

Robustness

☐  KRP-37 Rendering ten reports in sequence causes no OutOfMemoryError on a 3 GB device.

☐  KRP-38 NaN in a sample list does not crash the renderer.

☐  KRP-39 A render failure is logged under tag ECGReportScreen and does not crash the app.

☐  KRP-40 generateECGPdf returning null surfaces *"PDF generation failed"* to the user.

# 17. Open items

Ordered by impact. None is a crash today; the first two are the ones a user would notice.

| # | Item | Section |
| --- | --- | --- |
| 1 | isExporting never set — Download has no spinner and no double-tap guard | §10.1 |
| 2 | 2×6 pairs augmented leads as aVR, aVF, aVL rather than aVR, aVL, aVF — confirm intent | §7.3 |
| 3 | 41 MB ARGB_8888 bitmap on a 1080 px device; OOM risk on low-RAM hardware | §10.2 |
| 4 | Pan is unbounded — the report can be flung off-screen irrecoverably | §10.2 |
| 5 | shared_pdfs/ grows without cleanup on every Share | §11 |
| 6 | Locale.getDefault() in filename timestamps breaks backend filename parsing on non-Gregorian locales | §11 |
| 7 | isForPdf inferred from pxPerMm — misfires in small/split-screen windows | §5.2 |
| 8 | No decimation: up to 42,000 drawLine calls per report | §9 |
| 9 | Conclusion overflow silently truncates, and the two orientations number the survivors differently | §7.5 |
| 10 | adcScaleFactor is computed and never used | §13 |
| 11 | Duplicate stale copy at test_cloud_connection/ReportScreen.kt | §12 |
| 12 | Margin comments say 10 mm; the constants are 5 mm | §4 |